## Slide 1: SNOLAB Ambient Fast Neutron Studies
Brianna Liz Binoy
Research Assistant, CUTE Supervisors: Dr. Matt Stukel

## Slide 2
Neutron background at SNOLAB
Being underground doesn’t mean that we’re free of all radiation
2

## Slide 3
Neutron Sources At SNOLAB
1. Calibration Sources
2. Surrounding Norite Rock & Shotcrete
Cf-252
(Alpha, n) reactions
Fission reactions
3

## Slide 4
NEUTRON DETECTION
IT IS IMPORTANT WE DETECT and understand THEM
4

## Slide 5
Importance of Neutron Detection At SNOLAB
1. Neutron signals similar to dark matter signals
2. Is the neutron background homogenous across the lab?
5

## Slide 6
6
Neutron classifications | Energy
Cold energies | < 0.01 eV
Thermal | 0.01 eV – 0.3 eV
Epithermal | 0.3 eV-10 keV
Fast | 10 keV-20 MeV
Relativistic | > 20MeV
Neutron classifications
Notes: Spectrum

## Slide 7
NEUTRON DETECTOR
The ARKTIS FAST & Thermal NEUTRON DETECTOR
7

## Slide 8: Arktis S670e Detector
Made of stainless steel
Sensitive to fast and thermal neutrons and gamma rays
Fast neutrons: Pressurized Helium-4 gas & Thermal neutrons: Lithium-6 coating
    3 optically decoupled segments, each with 8 Silicon Photomultipliers
Arktis S670e Operation Manual
Notes: Talk to the slide more, point to the detector and show exactly where things are

## Slide 9: How the neutron signals are detected
Detects a signal
Measures how long the signal amplitude stays above a certain offset (the “threshold”)
If above threshold:
-If the signal amplitude is tall and wide: a neutron.
-If it’s small and quick: noise or a gamma ray
The Time over Threshold (ToT) parameter
X axis: time (s), y axis: signal amplitude (V)

## Slide 10
MY PROJECT
10
PART 1 | GET on-surface FAST neutron flux
PART 2 | GET UNDERGROUND FAST NEUTRON FLUX
PART 3 | UNDERSTAND THE DETECTOR’S NEUTRON SENSITIVITY

## Slide 11
FAST NEUTRON FLUX
Part 1: obtaining on-surface FAST neutron flux
11

## Slide 12
Method to get on-surface fast neutron flux
Placed Neutron detector in Lab B of the on-surface cleanroom
Took background data
Co60 gamma source data
Estimated on-surface fast neutron flux
Detector taking data in on-surface cleanroom
Notes: pictures

## Slide 13
Gamma counts
Fast neutron counts
Thermal neutron counts
Background Spectrum
Notes: Error bars

## Slide 14
2. Co-60 spectrum
Co60

## Slide 15
Gamma counts
DETECTED FAST NEUTRONS: 6 ± 2 fast neutrons/cm²/day
Fast neutron counts
Background flux and gamma flux merging
3. Obtaining fast neutrons detected by detector
cut obtained from literature (Browne, 1999)
Thermal
neutron
counts

## Slide 16
4. Applying intrinsic detection efficiency
Simulated the detector
Sent fast neutrons of different energies into detector simulation
Counted neutrons that entered detector & counted He4 reactions in detector
Intrinsic efficiency: (reactions counts/ neutron entries)%
DETECTOR data:
6 ± 2 fast neutrons/cm²/day
Data points
On – Surface Fast Neutron Flux neutrons/cm2/day | EXPACS Simulation Fast Neutron Flux neutrons/cm2/day
109 - 300 | 185
Notes: Make the numbers very big
title

## Slide 17
FAST NEUTRON FLUX
PART 2: Obtain underground Fast neutron flux
17

## Slide 18
Locations
J-drift
Ladder Lab B
SNOLAB' underground map
Purpose | Location | Shielding
Take background data | J-Drift | None
Take background data | Ladder Lab B | None
Taking background data in the J-Drift

## Slide 19
Underground flux against Surface flux
Fast neutron count region

## Slide 20
Obtaining Underground Fast Neutron Flux
Location | Detected Fast Neutron Flux (neutrons/cm2/day)
J-Drift | 3.57 ± 0.02
Ladder Lab | 3.12 ± 0.02
Data points
Fast Neutron Flux (neutrons/cm2/day)
67 - 179
56 - 157
On surface | 6 ± 2
109 - 300

## Slide 21
NEUTRON DETECTOR
PART 3: Understanding detector Neutron sensitivity
21

## Slide 22
Location
Low Background Lab (LBL)
SNOLAB' underground map
Purpose | Location | Neutron Shielding
Understand detector limits | Low Background Lab | Polyethylene & UPW
UPW & Polyethylene shielding of detector in LBL

## Slide 23
Shielding & No Shielding give similar neutron flux data
Location | Detector Data (fast neutrons/cm2/day)
No neutron shielding | 3.12 ± 0.02
LBL with neutron shielding | 2.7 ± 0.1
Spectra of shielded and unshielded data
neutron count region

## Slide 24
Bombarding the detector with neutrons
Source | Neutrons emitted by source (neutrons/day) | Theoretical calculation (neutrons/day) | Difference
1 Cf252 | 19680x10⁵ ± 3x10⁵ | 19400 × 10⁵ | ~1.4%
2 Cf252 | 40910x10⁵ ± 8x10⁵ | 40300 × 10⁵ | ~1.5%
1 Cf-252 source in the middle of detector
2 identical Cf-252 sources in the
middle of the detector
Linear relationship as expected
GEANT4 simulation of setup
Spectra of Cf252 sources

## Slide 25
Summary
Created Geant4 simulation of neutron detector
Got ON-SURFACE fast neutron flux range: 109 – 300 neutrons/cm2/day
Got UG fast neutron flux range in various locations: 56 - 179 neutrons/cm2/day
Understood detector's neutron sensitivity: minimum limit of 2.7 ± 0.1 neutrons/cm2/day
Next Steps
Get a value for the fluxes, rather than ranges
Look to see if there is any relationship between ToT values and neutron energies
Place detector outside to get surface flux
Notes: Cocnluding slide

## Slide 26
Thank you:
Matthew Stukel, Kishan Chaudhary, Ian Lawson, for helping me with my research
And also:
Connor Stubbs, Peter Qin & Qusai Almasri, for driving me to & from SNOLAB & around the extremely car-friendly Sudbury when I needed it
🌸
🌸
🌸

## Slide 27
Properties of Californium-252
Specific Activity: 0.536 mCi/μg (or 536 μCi/μg)
Neutron Yield: 2.314 × 10⁶ neutrons/second per μg
Activity in μCi: 5.2 μCi = 0.0052 mCi
Mass equivalent: 0.0052 mCi / 0.536 mCi/μg ≈ 0.0097 μg (9.7 ng)
Neutrons per second:
Neutrons per day: 22,445 n/s × 86,400 seconds/day ≈ 1.94 × 10⁹

## Slide 28
NEUTRON DETECTOR
PART 3: Simulating cavern neutron flux
28

## Slide 29
Process:
Obtain rock composition [1]
Did calculations to get the fraction of each isotope of each element in rocks
- Obtain U-238 and Th-232 decay chain information [2] to calculate fraction of all the radioactive daughter nuclei in the rocks
 | (Total neutrons/cm3/year)
Norite Yield | 12.97
Shotcrete Yield | 3.07
Theoretical Neutron Spectrum from SOURCES4C

## Slide 30
23592 U
23190 Th
42 He
Reacts with elements like Carbon, oxygen and beryllium
10 n
(Alpha, n) reactions
Uranium in the rocks
30

## Slide 31
Neutron Sources At SNOLAB
(Alpha, n) reactions
Fission reactions
31
2. Surrounding Norite Rock & Shotcrete

## Slide 32
Fission reactions
23592 U
Lighter element
Uranium in the rocks
Lighter element
10 n
Neutron from other
reactions
10 n
10 n
10 n

## Slide 33
Fast neutron
He-4 nucleus in detector
Detector wall
How Arktis Detects FAST Neutrons
33

## Slide 34
Fast neutron
He-4 nucleus in detector
Detector wall
How Arktis Detects FAST Neutrons
34

## Slide 35
----------->
Fast neutron
------>
He-4 nucleus in detector
Detector wall
How Arktis Detects FAST Neutrons
35

## Slide 36
SiPM
Electronic board
produces pulses
Scintillating light
How Arktis Detects FAST Neutrons

## Slide 37
Thermal neutron
Li-6 nucleus in detector
Detector wall
How Arktis Detects THERMAL Neutrons
37

## Slide 38
Thermal neutron
Detector wall
Li-6 nucleus in detector
How Arktis Detects THERMAL Neutrons
38

## Slide 39
---------->
High-energy alpha particle
--------------->
Triton nucleus
Detector wall
How Arktis Detects THERMAL Neutrons
39

## Slide 40
Scintillating light
produces pulses
High-energy alpha particle
SiPM
Electronic board
How Arktis Detects THERMAL Neutrons
40

## Slide 41
Seg1
Seg2
Seg3
Co60
Estimating on-surface neutron flux
Notes: Error bars

## Slide 42
Seg1
Seg2
Seg3
Co60
Estimating on-surface neutron flux
Notes: Error bars

## Slide 43
Seg1
Seg2
Seg3
Co60
Estimating on-surface neutron flux
Notes: Error bars

## Slide 44
5. Cross-checking result with EXPACs simulation
WHAT IS EXPACS?
EXcel-based Program for calculating Atmospheric Cosmic-ray Spectrum
It is based on the PARMA model: PHITS based Analytical Radiation Model in the Atmosphere,
Comprises analytical functions with parameters fitted to reproduce the results of the extensive air shower simulations
Detector Fast Neutron Flux neutrons/cm2/day | EXPACS Fast Neutron Flux neutrons/cm2/day
109 - 300 | 185
Notes: Just show neutrons, muons instead of all of them

## Slide 45
Our experimental data VS literature
Looked at previous experimental fast neutron data from MEASUREMENT OF THE FLUX AND ENERGY SPECTRUM OF COSMIC-INDUCED NEUTRONS ON THE GROUND
Integrated the spectrum
Used my intrinsic efficiency simulation to get fast neutron flux detected by detector
Detector Fast Neutron Flux (neutrons/cm2/day) | EXPACS Fast Neutron Flux (neutrons/cm2/day) | Literature’s Fast Neutron Flux (neutrons/cm2/day)
6 ± 2.0 fast neutrons | 3.5 - 9 | 2 -7

## Slide 50
Time over Threshold Parameter
Measures how long the signal amplitude stays above a certain offset (the “threshold”)
If the signal amplitude is tall and wide—possibly by a neutron
If it’s small and quick—possibly just noise or a gamma ray
Cuts can be set to determine the characteristic of the neutrons
The Time over Threshold (ToT) parameter
X axis: time (s), y axis: signal amplitude (V)

## Slide 51: HOW THE SIGNALS ARE DETECTED
SIGNAL CAPTURE
Each SiPM pair sends its signal to a discriminator
THRESHOLD CHECK
If signal > offset threshold  line is set to HIGH in Coincidence Unit
COINCIDENCE REQUIREMENT
If 2 lines are high within 30 ns, ToT counter starts
DECISION
If ToT within the cuts  likely a neutron  pulse produced
If not  not a neutron  no signal sent

## Slide 52
NEUTRON DETECTOR
GEANT4 SIMULATIONS OF OUR NEW DETECTOR
52

## Slide 53
53
LiF coating
He-4 gas
GEANT4 Simulation of Detector
x
y
z
Notes: Maybe show the axis or remove them

## Slide 54: GEANT4 SIMULATIONS OF DETECTOR: alpha and neutrons
1MeV Alpha particles
World Material: Vacuum
LiF cylinder
25meV neutrons
World Material: Vacuum
LiF cylinder
He-4 cylinder
1MeV Alpha particles
World Material: Vacuum
1 MeV neutrons
World Material: Vacuum
He-4 cylinder

## Slide 55
Testing Neutron Detector Simulation
Energy deposited by different neutron energies
Deposited Energy (MeV)
Events
Notes: Practice what I want to say more, remove the title

## Slide 56
Neutron and 𝜸-ray Discrimination by a Pressurized Helium-4 Based Scintillation Detector
Testing Neutron Detector Simulation
Energy deposited by different neutron energies
Deposited Energy (MeV)
Events
Notes: point out similarities in the both of them

## Slide 57
Testing Neutron Detector Efficiency
Fast neutron dose rate monitoring using off the shelf Helium-4 scintillation detectors
EXPERIMENT
Notes: Paper done previously
Tested the detector by sending neutrons of 3 different energies
See that the intrinicis efficiency decreased with neutron energy
So I did the same with my simulation- tested the efficiency

## Slide 58
Testing Neutron Detector Efficiency
Fast neutron dose rate monitoring using off the shelf Helium-4 scintillation detectors
SIMULATION
EXPERIMENT
Notes: This is the plot I have
You can see that when I sent same energies of neutrons, I got the exact same plot
This probably indicates that the neutron detector has very high efficiency, and the noise in the detector is very low. But we will need to do more analysis to confirm this.

## Slide 59: Cryogenic Underground TEst facility
Cleanroom
Ultra-pure water
Dry dilution fridge
59
Notes: The full form of CUTE is Cryogenic undergroudn testing facitily and it is located 2 kilometers underground at SNOLAB. But CUTE isn’t a detector itself. Instead, it’s a testing ground for experiments like dark matter and qubit characterizations.
CUTE consists of a clean room, a dilution fridge, a big tank of ultra pure water to shield components from neutron radiation
Because it’s so far underground, CUTE is shielded from cosmic radiation. And to test detectors that are supposed to catch very rare signals, we need to make sure the environment is as “quiet” as possible—no noise, no heat, and no shaking. CUTE uses a dilution refrigerator, clever trick with helium-3 and helium-4 gases to reach mind-blowing temperatures—down to 12 millikelvin, or about -273°C.
Formerly used for testing SuperCDMS dark matter detectors, CUTE now examines quantum chip performance in the absence of cosmic radiation.

## Slide 61
PREVIOUS NEUTRON MEASUREMENTS
61

## Slide 62
History of Measuring The Ambient Fast Neutron Flux
1. Found elements in norite rock and shotcrete
2. Use SOURCES4C and GEANT4
Notes: Estimated fast neutron flux: 4000 neutrons/m2/day (without experimental verification which is the problem)

## Slide 63
ISOTOPE | NORITE |  | SHOTCRETE | 
 | Concentration (ppm) | Neutron Flux (neutrons/year/cm3) | Concentration (ppm) | Neutron Flux (neutrons/year/cm3)
Total |  | 12.83 |  | 3.07
NORITE YIELD
SHOTCRETE YIELD
GEANT4: Estimated fast neutron flux of 4000 neutrons/m2/day
SOURCES4C Plots & Table
Notes: List where you got this information (maybe a source/name of paper)

## Slide 64
MY WORK
MY WORK TO OBTAIN THE ESTIMATED FAST NEUTRON SPECTRUM
64

## Slide 65
Tried to replicate their work
1. Obtain rock & shotcrete composition
2. Obtain U-238 and Th-232 decay chain information
Notes: Estimated fast neutron flux: 4000 neutrons/m2/day (without experimental verification which is the problem)

## Slide 66: New Theoretical Fast Neutron Spectrum from SOURCES4C
NORITE YIELD
SHOTCRETE YIELD
 | Old Data (Total neutrons/cm3/year) | My Data (Total neutrons/cm3/year)
Norite Yield | 12.83 | 12.97
Shotcrete Yield | 3.07 | 8.70
Notes: A lot of info on the slides
Remove the references
Speak about the process instead , put the title, then show the slides / break into 2 slides

## Slide 67
67
Arktis Discrimination Parameter
Arktis S670e Operation Manual
Notes: Counts as a function of the Time-over-Threshold discrimination parameter. The blue line is
the counts from an un-moderated Cf-252 source; the red one is the counts from a moderated Cf-252
source; and the green one is the pure gamma counts. The cuts to discriminate between the
gammas, fast and thermal neutrons are based on this spectrum.

## Slide 68
Fast neutron
He-4 nucleus in detector
How Arktis detects FAST neutrons
SiPM
Electronic board
The S-670e produces TTL pulses of different widths
68